UltraScan Workshop/Bioc5083
Hydrodynamic Methods
Borries Demeler, Ph.D.
Department of Biochemistry
May/June 2013

---
Analytical Ultracentrifugation
Background
What can be learned from AUC?
Excellent method for characterizing any molecule or molecular interaction in the solution environment – small sample requirement, up to 14 samples can be analyzed. Analysis is based on first principles - No standards are required
Molecules can be studied in a physiological environment – solution conditions can be adjusted (concentration dependency, effect of pH, ionic strength, buffer type, ligands, oxidation state, temperature, etc.)
Very large size range (102 – 108 Dalton)
Dynamics - measure oligomerization states of reversible self- or hetero-associations, ligand binding, slow kinetics and Kd
Composition analysis – number of components, their partial concentration, molecular weight, and anisotropy
Conformational analysis - folding/melting studies of biopolymers, conformational changes based on changes in solution properties

---
Analytical Ultracentrifugation
Background
Available Optical Systems:
Multiple detectors extend the range of AUC applications:
UV/vis absorbance (single wavelength): all-purpose detector for biopolymers and materials absorbing in the visible. Slower detection, good for low protein concentrations.
Fluorescence: Study molecules with intrinsic fluorophores or eGFP fusions in impure cell extracts, exquisite selectivity for binding experiments, add fluorescently labeled antibodies to an impure cell extract, study the order of assembly in a multi-domain protein complex
Interference: fast data acquisition and measurement of non-absorbing molecules, carbohydrates, high-concentration studies
Multi-wavelength UV/vis: obtain additional spectral dimension in addition to hydrodynamic separation for independent characterization of molecular properties

---
AUC Instrumentation
UV/vis absorbance, Rayleigh interference, fluorescence
Multiple recordings over time...

---
Sedimentation Review
Sedimentation:
Forces at Equilibrium:
Fc - Fb - Fd = 0
Fb (buoyancy)	= ω2 r ms
Fd (viscous drag)	= fv
Fc (centrifugal force)	= ω2 r m
Explanation:
Fb is the buoyancy force - the force	required to displace the buffer surrounding	the solute, and ms is the mass of the	displaced solvent.

---
Fb (buoyancy)	= ω2 r ms
Fd (viscous drag)	= f v
Fc (centrifugal force)	= ω2 r m
Substitute the mass of the solvent, m0,
with the mass of the solute, m
Rearrange the force equation:
Fc - Fb - Fd = 0 and substitute
Place molecular parameter on one side
and experimental parameters on the other
Put into molar units by multiplying with
Avogadro's number, N
m_s~ = ~ m bar v %rho , ~Fb ~= ~ %omega^2 r m bar v %rho
%omega^2 r m ~ - ~ %omega^2 r m bar v %rho ~ = ~fv
{m`left (1`-` bar v %rho` right )} over f ~ = ~ v over {%omega^2 r}
{M`left (1`-` bar v %rho` right )} over {Nf} ~ = ~ v over {%omega^2 r} ~= ~ s

---
Partial Specific Volume and Buoyancy
Sedimentation Forces:	Fc - Fb - Fd = 0
Fb (buoyancy) = 2 r m0
Fd (viscous drag) = fv
Fc (centrifugal force) = 2 r m
We need:
Partial specific volume ( ) of solute (we can estimate this from the protein sequence or measure in a densitometer – UltraScan will estimate from protein sequence automatically)
Viscosity of solvent (we can obtain this measure from the known composition of the buffer or measure in a viscometer – UltraScan will estimate this quantity from buffer composition automatically)
Density of solvent (we can obtain this measure from the known composition of the buffer or measure it in a densitometer – UltraScan will estimate this quantity from buffer composition automatically)
Temperature of experiment (recorded by instrument)
Rotor speed (recorded by instrument)
bar %nu
bar %nu
{M`left (1`-` bar v %rho` right )} over {N`f} ~ = ~ v over {%omega^2 r} ~= ~ s

---
Definition:
The sedimentation velocity, v, divided by the centrifugal field strength, ω2r, is equal to the sedimentation coefficient, s
Take-home message:
The sedimentation coefficient is proportional to M and inversely proportional to f
{M`left (1`-` bar v %rho` right )} over {N`f} ~ = ~ v over {%omega^2 r} ~= ~ s
[Speaker Notes: Take home message: The sedimentation coefficient is directly proportional to the molecular weight of the solute, and inversely proportional to the frictional coefficient of the solute. A large molecular weight increases the sedimentation speed, while an irregular shape will slow it down.]

---
Radial Dilution
Radial Dilution occurs because of the cell's sector shape. Molecules sedimenting towards the outside of the cell will dilute as they sediment.
All molecules - no matter at what position they are - will dilute at the same rate, causing a reduction in the observed optical density. At any given time, this dilution is the same at each point in the cell.
Radial Dilution can be observed through a reduction in the plateau absorbance in successive scans
[Speaker Notes: Due to the sector shape of the analytical ultracentrifugation cell, molecules will dilute as a consequence of transport to a larger volume element during sedimentation. The graphic in the upper left schematically shows a lamella of molecules at the inside of the cell sedimenting towards the outside of the cell, and diluting in the process. The graphic in the lower left shows the effect of radial dilution on the plateau concentration which successively decreases with each time point.]

---
Effect of Diffusion on the Sedimentation Boundary
[Speaker Notes: The effect of diffusion: Instead of a step function we observe a flattening of the boundary that increases over time. If the sample had no diffusion, transport due to sedimentation would proceed as a step function, and the moving boundary would be a straight line with an infinite gradient (blue line). All but the very largest molecules exhibit measurable diffusion over the time course of a velocity experiment, resulting in a broadening of the moving sedimentation boundary over time (indicated by the red line). This results in boundaries becoming broader over time.]

---
Transport Processes – Diffusion
Random translation	of a particle due to	Brownian motion
The flow due to	diffusion is	proportional to the	concentration	gradient and the	diffusion coefficient	(Fick's first law):
The rate of change of concentration is proportional to the change in steepness of concentration gradient (Fick's second law):
D~=~RT over {N color red f}
J~=~-D {partial C} over {partial x}
{partial C} over {partial t}~=~D {partial^2 C} over {partial x^2}
[Speaker Notes: Particles suspended in a solution experience Brownian motion, a random translocation of the particle that increases with increasing temperature. This translocation is quantified by the translational diffusion coefficient. The translational diffusion coefficient, D, is proportional to the temperature and inversely proportional to the shape, or the frictional coefficient f . This makes the diffusion transport interesting to AUC. It provides a second measure for the shape. Like sedimentation, diffusion is also inversely proportional to shape. If we know s and D, we can calculate the molecular weight (using Sverdberg's equation). The flow J is proportional to the diffusion coefficient and also to the steepness of the concentration gradient. The steeper the gradient gets, the stronger the propensity to “equilibrate” the gradient to a less steep gradient. In a moving boundary (essentially unrestricted on either side until the meniscus or bottom of cell are encountered), the solute can diffusion ahead of the boundary or back away from the boundary, so in the center or midpoint of the boundary the solute doesn't know which way to flow, consequently there is no diffusion occurring at the midpoint. This is explained by Fick's first law. The diffusion speed (translation with respect to time) is described by the derivative of the concentration gradient, which says that at the midpoint the diffusion speed is zero, and at the lagging end of the boundary the speed is largest half-way between the baseline and the midpoint, and in the direction away from the boundary movement, and above the midpoint the diffusion speed is greatest half-way between the midpoint and the plateau and in the direction of the sedimenting boundary.]

---
Calculation of the Sedimentation Coefficient:
int from r=m to r=b {1 over r} dr ~=~ int from t=0 to t=scan s %omega^2 dt
{M`left (1`-` bar v %rho` right )} over {N`f} ~ = ~ v over {%omega^2 r} ~= ~ s
hat s `_b ~=~ ln {left ( {r_b`(t)} over {r_m ` (t_0)} right )} ~ left (%omega^2`(t`-t_0) right )^{`-1}
v ~= ~ dr over dt
dr over r ~= %omega^2 ` s ` dt
[Speaker Notes: If the apparent sedimentation coefficient is calculated from different boundary positions as shown in the previous slide, the resulting sedimentation coefficient is not constant over time, unless the midpoint boundary position is chosen. This again is a result of diffusion, which causes the apparent sedimentation coefficient to decrease over time if the boundary position is above the midpoint, and to increase if the boundary position is below the midpoint. Clearly, extrapolation to infinite time will cause all boundary positions to converge to the sedimentation coefficient from the midpoint. It may be counterintuitive that the deviation of s-values is greatest near the beginning of the experiment, since clearly diffusion has not had a lot of time to affect transport. The explanation is that the concentration gradient is largest at the beginning of the experiment, causing the fastest transport due to diffusion. Another observation is that the deviation for boundary positions above and below the midpoint are not linear with time. The van Holde – Weischet method addresses both these observations, as will be seen in the following slides.]

---
Viscosity and Density Corrections
The density and also the viscosity of the solvent affect the sedimentation and diffusion of the particle in solution, so the measured values need to be corrected to standard conditions. Moreover, temperature and buffer composition affect the solvent
density and viscosity, so they need to be considered. To correct
for density and viscosity, use these formulas:
s_{20,W} ~=~ s_{T,B} {left ( 1- bar %nu %rho right )_{20,W}` %eta_{T,B}} over {left ( 1- bar %nu %rho right )_{T,B}` %eta_{20,W}}
D_{20,W} ~=~ D_{T,B} {293.15 ` %eta_{T,B}} over {T ` %eta_{20,W}}
UltraScan will automatically apply these corrections for aqueous buffers!

---
Transport Processes – Sedimentation and Diffusion
At Rest
rotor speed = 0
[Speaker Notes: At rest, all solutes are distributed equally throughout the cell and the absorbance is uniform.]

---
Transport Processes – Sedimentation and Diffusion
Sedimentation Velocity
        Duration is hours
        high rotorspeed
At Rest
rotor speed = 0
[Speaker Notes: A moving boundary that diffuses is created when the rotor is accelerated. This creates an absorbance pattern where the meniscus region is depleted from absorbing molecules over time.]

---
Transport Processes – Sedimentation and Diffusion
left ( {partial C} over {partial t} right ) `_r ~ = ~ { - `1} over r partial over {partial r}~ left [` s %omega^2 r^2 C~-~ D r` {{partial C} over {partial r } }` right ] `_t
Sedimentation Velocity
        Duration is hours
        high rotorspeed
At Rest
rotor speed = 0
[Speaker Notes: The Lamm equation models the flow of the sedimenting solute in a sector shaped cell that has a boundary condition (meniscus and bottom of cell constraint). It models both the sedimentation and the diffusion, and accounts for the diffusion resulting from the concentration gradient at the boundary and at the bottom of the cell.]

---
Transport Processes – Sedimentation and Diffusion
left ( {partial C} over {partial t} right ) `_r ~ = ~ { - `1} over r partial over {partial r}~ left [` s %omega^2 r^2 C~-~ D r` {{partial C} over {partial r } }` right ] `_t
J `=` s %omega^2 r C~-` D {partial C} over {partial r} ~=` 0
Sedimentation Velocity
        Duration is hours
        high rotorspeed
Sedimentation Equilibrium
        Duration is > 1 day
        low rotorspeed
At Rest
rotor speed = 0
[Speaker Notes: For equilibrium experiments, the net flow in the cell is zero at equilibrium, which avoids the effect of friction on the distribution of the solute. At equilibrium, an exponential gradient is formed that is proportional to the weight-average molecular weight of all solutes in the sample.]

---
Relationship between M, f, k, s, and D
matrix {alignl {color red {bold f (D)}} # { ~=~ } # alignl {{{RT} over {N color red {bold D}}}} ## { } # { } # { } ## alignl {color red { bold M (s, D)}} # { ~=~ } # alignl {{{color red { bold s} RT} over {color red { bold D }left ( 1 `-` bar %nu %rho right)} }} ## { } # { } # { } ## alignl {color red {bold f_0 (M)}} # { ~=~ } # {%eta left ( {162 color red { bold M} %pi^2 bar %nu} over {N} right )^{1/3} ~~~~~r_s ~=~ color red f_0 over {6 %pi %eta} ~~~~~~~~%phi~=~ f over f_0} ## { } # { } # { } ## alignl {color red {bold D (s, %phi)}} # { ~=~ } # alignl { {RT left [ N 18 %pi left (color red { bold %phi} %eta right )^{3/2} left ( {{color red { bold s} bar %nu } over {2 (1 - bar %nu %rho)}} right)^{1/2} right ]^-1}} }
left ( {partial C} over {partial t} right ) `_r ~ = ~ { - `1} over r partial over {partial r}~ left [` s %omega^2 r^2 C~-~ D r` {{partial C} over {partial r } }` right ] `_t
Concentration Sedimentation Diffusion
[Speaker Notes: The Lamm equation shown above is has a sedimentation term (blue) and a diffusion term (red). This means we need two observables to represent fully the solution behavior of a solute in a sedimentation experiment. Using the Stokes-Einstein relationship and the Svedberg equation, we can represent the diffusion coefficient as a function of s and frictional ratio, using the Svedberg equation alone, we can express molecular weight as a function of s and D. So the hydrodynamic solution properties of solutes in a sedimentation experiment can be represented in a variety of different parameterizations, each of which can be used for modeling the Lamm equation. For finite element fitting, we use the frictional ratio and sedimentation, since it is easy to come up with reasonable limits: The sedimentation limits can be obtained from the model independent van Holde – Weischet analysis, and the frictional ratio limits can be reasonably guessed from prior knowledge]

---
Parametrization of Anisotropy (Shape)
The frictional ratio f/f0 = k is a convenient way to parameterize the diffusion coefficient and the shape of a molecule .
The frictional ratio k is 1.0 for a sphere since f = f0 and hence k has a convenient lower limit
k = 1.0
1.0 ≤ k ≤ 4.0 for most proteins, higher for rod-shaped, disordered and unfolded proteins, DNA, fibrils and aggregates or linear molecules
k = 1.2 - 2.5
 k > 3
[Speaker Notes: f/f0 is a convenient parameterization for the diffusion coefficient, giving reasonable values between 1-10]

---
Degeneracy of Shape Determination
 k ( ) = k ( )
k may be the same for different shapes, we cannot distinguish them by AUC, we can only measure the anisotropy!
[Speaker Notes: Molecules can assume different shapes. This slides compares f/f0, s, and D for a molecule with the same molecular weight and density but with different conformations. f/f0 is a measure of globularity, if close to one, the molecule is spherical, if f/f0 is large, the molecule is in a random coil or rod-shaped conformation.]

---
UltraScan Layout (multi-platform, open source)
User
Web Server
MySQL DB
US LIMS
GridControl
UTHSCSA/XSEDE
UT Data Center
Apache/Airavata
Middleware
High Performance Computing Clusters
[Speaker Notes: The user interacts with the XLA, the UltraScan GUI program, which is written in platform independent C++, and the web server running on Jacinto. The user collects the data on the XLA/I/F and transfers it to UltraScan, which is used to store it in the MySQL database. The user can then go to a website and log into the USLIMS program, a PHP suite of modules that control access to the data through a series of web applications that control data sharing, data display, and analysis submission. Analysis jobs are submitted to the TIGRE/Globus grid of supercomputers. Such computers can be anywhere in the world and can be configured to share jobs for higher performance analysis.]

---
UltraScan HPC Sites
Juropa@Jülich
Godzilla @ Allergan
 IU
FU
Lonestar @ TACC
SD
Stampede@TACC
Trestles@SDSC
AU
Alamo@UTHSCSA
SA
Jacinto@UTHSCSA

---
 UltraScan Components
The User interacts with 3 UltraScan Modules:
1. UltraScan GUI:
Used for data upload, editing, desktop analysis, visualization, data management
2. UltraScan HPC:
High-Performance analysis module for 2-dimensional spectrum analysis, genetic algorithms, and Monte Carlo analysis
3. UltraScan LIMS:
Provides convenient data exchange
Assists with AUC multi-user facility management
Provides central data storage for experimental data and descriptions, peptide sequences, buffer conditions, investigator information, experimental conditions and status (tracking), spectra, gel images.
Provides web interface for investigators and analysis reports
Provides UltraScan interface for analysts
Provides web portal for supercomputer analysis

---
UltraScan Data Flow
Experimental Data, Buffers,
Peptide sequences, Designs
2
PC for Data
Acquisition
AUC
UltraScan
LIMS Server
MySQL
Database
UltraScan
LIMS Portal
(Apache Web
Server)
PC for Data
Editing and
Analysis
Experimental Data, Buffers,
Peptide sequences
3
Edited and
Analyzed Data
Genetic Algorithm
2D Spectrum Anal.
Monte Carlo Anal.
HPC
System
4
6
5
User/
Collaborator
7
Results
Project Description, Peptide sequences, Buffers, Gel Images, Absorption scans, Experimental requests
1
[Speaker Notes: Flow chart for the UltraScan Laboratory Information Management System (LIMS): 1. The user or collaborator submits a project request, describing the background of the planned experiment. Also submitted are supporting information such as peptide sequences, buffer composition, gel images, absorbance scans, etc. 2. The investigator prepares an experimental design and appends it to the notes in the project request and performs the experiment. The experimental data are then associated with the appropriate buffer and peptide files and committed to the database. 3. The analyst retrieves the experimental data, edits it and performs a preliminary van Holde – Weischet analysis to obtain s-ranges. 4. The edited data are committed to the database. 5. The data are now analyzed on the HPC system, and the results are sent by e-mail to the analyst. 6.The analyst can visualize the results and prepare an analysis report, which is uploaded along with the data to the LIMS 7. The collaborator can log into the web portal and review the analysis results]

---
DEMONSTRATION
UltraScan Software Demo:
Configuration of UltraScan-III and Database
Calibration of rotor stretch
Import of legacy data into UltraScan-III
Editing of velocity and equilibrium data

---
Modeling of Data
A first principles approach allows us to model the experimental data by fitting it to a mathematical model. The model represents the physics of the experiment, and contains parameters of interest to the experimentalist. We need to find the values of these parameters by adjusting the model so it matches the data. This is a hard problem called an “inverse problem” that requires optimization (fitting) algorithms which aid us in adjusting the parameters so the model fits the data.
In UltraScan this is accomplished by a least squares fitting approach that compares each data point from the model with the corresponding point in the experimental data:
Optimally, the difference is zero, but because of experimental noise this never happens, since the model is noise free.
Minimize SUM from i=1 to N left ( Data_i ~-~ Model_i right )^2 ~~"(" i " over radius and time)"

---
Noise & Data Modeling Considerations
Fitting of noisy data prevents unique solutions – multiple solutions are possible.
We need to minimize noise when modeling data.
There are three ways to reduce or eliminate noise:
1. fit the noise
2. maintain an exceptionally well tuned instrument
3. design your experiment to optimize the quality of the data
There are three noise types:
1. Time Invariant noise: Noise is different for each radial position, but the same offset for each scan, and hence time independent (finger prints).
2. Radially Invariant noise: Noise is different for each scan, but each radial position is offset by the same amount throughout the scan (baseline variation)
3. Stochastic (random noise): Noise is different for each radial and time point and it is (hopefully) Gaussian in distribution
(1) and (2) can be fitted by UltraScan and removed from the data

---
Different Types of Noise
Time Invariant noise: Noise is different for each radial position, but the same offset for each scan, and hence time independent.

---
Different Types of Noise
Radially Invariant noise: Noise is different for each scan, but each radial position is offset by the same amount throughout the scan

---
Different Types of Noise
Stochastic (random noise): Noise is different for each radial and time point and it is (hopefully) Gaussian in distribution:

---
Intensity vs. Absorbance
Time invariant noise sources:
Reference Channel:
Lamp window
Monochromator optics
Cell window (reference channel)
Slit assembly
Photomultiplier window
Sample Channel:
Lamp window
Monochromator optics
Cell window (sample channel)
Slit assembly
Photomultiplier window
Radially invariant noise sources:
Optical path length changes
Low frequency intensity changes
Stochastic noise sources:
Electronics
Photomultiplier tube
Lamp flashes

---
Tale of 2 noisy vectors
Noise of Scan 1
+
Noise of Scan 2
=
Noise *
bold sqrt 2

---
Intensity vs. Absorbance
  Optical system considerations
Intensity measurements record the intensity of light passing through one channel. Absorbance measurements record the intensity of light passing through one channel, then record the intensity of the light passing through the reference channel, and subtract it from the first channel. Each channel recording contains the (nearly) same amount of time invariant noise, but different amount of stochastic noise. Subtraction of time invariant noise will eliminate it:
scan_1 ~=~ signal_1` +` N_s1 `+` N_ti newline scan_2 ~=~ signal_2` +` N_s2 `+` N_ti newline scan_1 `–` scan_2 ~=~ signal_1 `–` signal_2 `+` N_s1 `+` N_s2 newline newline N_s1 `+` N_s2~ approx~ N_s1 `sqrt(2)

---
Advantages of Intensity Data
Measurement of intensity data when measuring velocity experiments has multiple advantages:
Stochastic noise is reduced by approximately a factor of 1.4, providing much better quality data and improving the RMSD of finite element solution data fits.
The capacity of the instrument is increased because the reference channel can be filled with sample, although the total OD in the reference cell should remain below 0.5 (for the reference cell). The sample OD can be higher, but just as in absorbance optics, it should not exceed the dynamic range of the UV-vis detector.
Due to the lower stochastic noise the sample concentration can be reduced while maintaining the same signal-to-noise ratio as observed in the absorbance measurement.

---
Advantages of Intensity Data
Additional considerations when measuring intensity data:
Intensity data can only be used for velocity experiments, since considerable time invariant noise and also increased radially invariant noise (from variations in the lamp intensity) is present in such data. Equilibrium data by definition are time-invariant and hence correlate 100% with the noise removal routine which is essential for using time-invariant noise.
A Reference Channel is no longer required, the air-to-air region above the meniscus can be used as a reference.
Baseline absorbance must be added to sample absorbance and the total must be lower than 0.5 OD to avoid gain setting changes of the photomultiplier tube.

---
Summary:
Time- and radially- invariant noise can be fitted by UltraScan and removed from the data to improve the results.
Stochastic noise cannot be removed and should be minimized by maintaining a well calibrated instrument and performing a well-designed experiment!
Data subtraction in absorbance mode convolutes two stochastic vectors and leads to an increase in stochastic noise by ~
Remember:
you cannot get reliable answers if you start
with low quality input data!
bold sqrt 2